## LMDADMINSMD-E-
DOC #: BRCC/MGT/FRM-053 ISSUE STATUS: 01 ISSUE DATE: Jan 24, 2018
In House Calibration Records
Sr. | Department | MACHINE NAME | Instrument Name | Calibration Code | Make / Model | Range/Least Count | Location | Calibration Date | Next Calibration Due Date | Deviation limit | Calibration Results |  |  | Remarks
[11 empty] | Instrument under calibration | Master Standard Value | Deviation
1 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-01-TG1 | 82 | CHINA | 1 | LMD-PRD | 2021-12-24 | 2022-12-24 | 5 C | 20 | 20 | 0 | Okay
[11 empty] | 40 | 40.1 | -0.1 | Okay
[11 empty] | 60 | 60.3 | -0.3 | Okay
2 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-01-PG1 | 81 | CHINA | 0.01 | LMD-PRD | 2021-12-24 | 2022-12-24 | 0.05 MPA | 0.5 | 0.53 | -0.03 | Okay
[11 empty] | 0.6 | 0.65 | -0.05 | Okay
[11 empty] | 0.6 | 0.65 | -0.05 | Okay
3 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-01-PG2 | 254 | CHINA | 0.01 | LMD-PRD | 2021-12-24 | 2022-12-24 | 0.05 MPA | 0.5 | 0.65 | -0.15 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.6 | 0.65 | -0.05 | Okay
4 | LMD-PRD | BEAD MILL | CCD-LMD-PRD-EQT-01-PG3 | 255 | CHINA | 0.01 | LMD-PRD | 2021-12-24 | 2022-12-24 | 0.05 MPA | 0.5 | 0.65 | -0.15 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.6 | 0.65 | -0.05 | Okay
5 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-01-AMP M01 | 257 | CHINA | 1 | LMD-PRD | 2021-12-24 | 2022-12-24 | 1A | 0.4 | 0.5 | -0.1 | Okay
[11 empty] | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
6 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-01-TIMER | 256 | CHINA | 1 | LMD-PRD | 2021-12-24 | 2022-12-24 | 1 | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.6 | 0.5 | 0.1 | Okay
[11 empty] | 0.4 | 0.5 | -0.1 | Okay
7 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-02-RPM | 79 | CHINA | 10 | LMD-LAB | 2021-12-24 | 2022-12-24 | 5RPM | 98 | 100 | -2 | Okay
[11 empty] | 95 | 100 | -5 | Okay
[11 empty] | 101 | 100 | 1 | Okay
8 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-02-TG | 80 | CHINA | 1 | LMD-PRD | 2021-12-24 | 2022-12-24 | 5 C | 25 | 25 | 0 | Okay
[11 empty] | 45 | 45.1 | -0.1 | Okay
[11 empty] | 60 | 60.3 | -0.3 | Okay
9 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-02-PG1 | 253 | CHINA | 0.01 | LMD-PRD | 2021-12-24 | 2022-12-24 | 0.05 MPA | 0.4 | 0.64 | -0.24 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.6 | 0.65 | -0.05 | Okay
10 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-02-PG2 | 251-A | CHINA | 0.01 | LMD-PRD | 2021-12-24 | 2022-12-24 | 0.05 MPA | 0.5 | 0.65 | -0.15 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.6 | 0.65 | -0.05 | Okay
11 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-02-PG3 | 252 | CHINA | 0.01 | LMD-PRD | 2021-12-24 | 2022-12-24 | 0.05 MPA | 0.5 | 0.65 | -0.15 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.6 | 0.65 | -0.05 | Okay
12 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-02-AMP1 | 250-A | CHINA | 1 | LMD-PRD | 2021-12-24 | 2022-12-24 | 1A | 0.4 | 0.5 | -0.1 | Okay
[11 empty] | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
13 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-02-PS1 | 391 | CHINA | 0.01 | LMD-PRD | 2021-12-24 | 2022-12-24 | 0.05 MPA | 0.5 | 0.65 | -0.15 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.6 | 0.65 | -0.05 | Okay
14 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-03-TM1 | 77 | CHINA | 1 | LMD-PRD | 2021-12-24 | 2022-12-24 | 5 C | 20 | 20 | 0 | Okay
[11 empty] | 40 | 40.1 | -0.1 | Okay
[11 empty] | 60 | 60.3 | -0.3 | Okay
15 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-03-AMP M1 | 246-A | CHINA | 1 | LMD-PRD | 2021-12-24 | 2022-12-24 | 1A | 0.6 | 0.5 | 0.1 | Okay
[11 empty] | 0.4 | 0.5 | -0.1 | Okay
[11 empty] | 0.2 | 0.5 | -0.3 | Okay
16 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-03-TS1 | 78 | CHINA | ON/Off | LMD-PRD | 2021-12-24 | 2022-12-24 | ON/Off | on/off | on/off | #VALUE! | Okay
[11 empty] | on/off | on/off | #VALUE! | Okay
[11 empty] | on/off | on/off | #VALUE! | Okay
17 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-03-PG1 | 247-A | CHINA | 0.01 | LMD-PRD | 2021-12-24 | 2022-12-24 | 0.05 MPA | 0.6 | 0.65 | -0.05 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.4 | 0.65 | -0.25 | Okay
18 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-03-PG2 | 248-A | CHINA | 0.01 | LMD-PRD | 2021-12-24 | 2022-12-24 | 0.05 MPA | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.4 | 0.65 | -0.25 | Okay
[11 empty] | 0.6 | 0.65 | -0.05 | Okay
19 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-03-PG3 | 249-A | CHINA | 0.01 | LMD-PRD | 2021-12-24 | 2022-12-24 | 0.05 MPA | 0.6 | 0.65 | -0.05 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.4 | 0.65 | -0.25 | Okay
20 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-04-TM1 | 75 | CHINA | 1 | LMD-PRD | 2021-12-24 | 2022-12-24 | 5 C | 10 | 8 | 2 | Okay
[11 empty] | 15 | 18 | -3 | Okay
[11 empty] | 15 | 18 | -3 | Okay
21 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-04-AMP1 | 242-A | CHINA | 1 | LMD-PRD | 2021-12-24 | 2022-12-24 | 1A | 0.4 | 0.5 | -0.1 | Okay
[11 empty] | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
22 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-04-TS1 | 76 | CHINA | ON/Off | LMD-PRD | 2021-12-24 | 2022-12-24 | ON/Off | on/off | on/off | #VALUE! | Okay
[11 empty] | on/off | on/off | #VALUE! | Okay
[11 empty] | on/off | on/off | #VALUE! | Okay
23 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-04-PG1 | 245-A | CHINA | 0.01 | LMD-PRD | 2021-12-24 | 2022-12-24 | 0.05 MPA | 0.6 | 0.65 | -0.05 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.4 | 0.65 | -0.25 | Okay
24 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-04-PG2 | 243-A | CHINA | 0.01 | LMD-PRD | 2021-12-24 | 2022-12-24 | 0.05 MPA | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.4 | 0.65 | -0.25 | Okay
[11 empty] | 0.6 | 0.65 | -0.05 | Okay
25 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-04-PG3 | 244-A | CHINA | 0.01 | LMD-PRD | 2021-12-24 | 2022-12-24 | 0.05 MPA | 0.6 | 0.65 | -0.05 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.4 | 0.65 | -0.25 | Okay
26 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-04-PS1 | 390 | CHINA | ON/Off | LMD-PRD | 2021-12-24 | 2022-12-24 | 0.05 MPA | 0.5 | 0.65 | -0.15 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.6 | 0.65 | -0.05 | Okay
27 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-05-PG4 | 239 | CHINA | 0.01 | LMD-PRD | 2021-12-24 | 2022-12-24 | 0.05 MPA | 0.4 | 0.65 | -0.25 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.6 | 0.65 | -0.05 | Okay
28 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-05-PG3 | 240 | CHINA | 0.01 | LMD-PRD | 2021-12-24 | 2022-12-24 | 0.05 MPA | 0.4 | 0.65 | -0.25 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.6 | 0.65 | -0.05 | Okay
29 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-05-PG1 | 242-B | CHINA | 0.01 | LMD-PRD | 2021-12-24 | 2022-12-24 | 0.05 MPA | 0.6 | 0.65 | -0.05 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.4 | 0.65 | -0.25 | Okay
30 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-05-PG2 | 241 | CHINA | 0.01 | LMD-PRD | 2021-12-24 | 2022-12-24 | 0.05 MPA | 0.4 | 0.65 | -0.25 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.6 | 0.65 | -0.05 | Okay
31 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-05-TC | 75 | CHINA | 1 | LMD-PRD | 2021-12-24 | 2022-12-24 | 5 C | 10 | 8 | 2 | Okay
[11 empty] | 15 | 18 | -3 | Okay
[11 empty] | 15 | 18 | -3 | Okay
32 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-07-TM1 | 83 | CHINA | 1 | LMD-PRD | 2021-12-24 | 2022-12-24 | 5 C | 25 | 25 | 0 | Okay
[11 empty] | 45 | 45.1 | -0.1 | Okay
[11 empty] | 60 | 60.3 | -0.3 | Okay
33 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-07-TS1 | 84-A | CHINA | ON/Off | LMD-PRD | 2021-12-24 | 2022-12-24 | ON/Off | on/off | on/off | #VALUE! | Okay
[11 empty] | on/off | on/off | #VALUE! | Okay
[11 empty] | on/off | on/off | #VALUE! | Okay
34 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-07-PG1 | 296-A | CHINA | 0.01 | LMD-PRD | 2021-12-24 | 2022-12-24 | 0.05 MPA | 0.6 | 0.65 | -0.05 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.4 | 0.65 | -0.25 | Okay
35 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-07-PG2 | 294-A | CHINA | 0.01 | LMD-PRD | 2021-12-24 | 2022-12-24 | 0.05 MPA | 0.4 | 0.65 | -0.25 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.6 | 0.65 | -0.05 | Okay
36 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-07-PG3 | 293-A | CHINA | 0.01 | LMD-PRD | 2021-12-24 | 2022-12-24 | 0.05 MPA | 0.6 | 0.65 | -0.05 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.4 | 0.65 | -0.25 | Okay
37 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-07-AMP1 | 292-A | CHINA | 1 | LMD-PRD | 2021-12-24 | 2022-12-24 | 1A | 0.2 | 0.5 | -0.3 | Okay
[11 empty] | 0.9 | 0.5 | 0.4 | Okay
[11 empty] | 0.1 | 0.5 | -0.4 | Okay
38 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-07-PS1 | 395 | CHINA | ON/Off | LMD-PRD | 2021-12-24 | 2022-12-24 | 0.05 MPA | 0.5 | 0.65 | -0.15 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.6 | 0.65 | -0.05 | Okay
39 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-06-TC1 | 85 | CHINA | 1 | LMD-PRD | 2021-12-24 | 2022-12-24 | 5 C | 10 | 8 | 2 | Okay
[11 empty] | 15 | 18 | -3 | Okay
[11 empty] | 15 | 18 | -3 | Okay
40 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-06-AMP1 | 400 | CHINA | 1 | LMD-PRD | 2021-12-24 | 2022-12-24 | 1A | 0.2 | 0.5 | -0.3 | Okay
[11 empty] | 0.9 | 0.5 | 0.4 | Okay
[11 empty] | 0.1 | 0.5 | -0.4 | Okay
41 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-06-TS1 | 84-B | CHINA | ON/Off | LMD-PRD | 2021-12-24 | 2022-12-24 | ON/Off | on/off | on/off | #VALUE! | Okay
[11 empty] | on/off | on/off | #VALUE! | Okay
[11 empty] | on/off | on/off | #VALUE! | Okay
42 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-06-PG1 | 397 | CHINA | 0.05 MPA | LMD-PRD | 2021-12-24 | 2022-12-24 | 0.05 MPA | 0.6 | 0.65 | -0.05 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.4 | 0.65 | -0.25 | Okay
43 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-06-PG2 | 398 | CHINA | 0.01 | LMD-PRD | 2021-12-24 | 2022-12-24 | 0.05 MPA | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.4 | 0.65 | -0.25 | Okay
[11 empty] | 0.6 | 0.65 | -0.05 | Okay
44 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-06-PG3 | 399 | CHINA | 0.01 | LMD-PRD | 2021-12-24 | 2022-12-24 | 0.05 MPA | 0.6 | 0.65 | -0.05 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.4 | 0.65 | -0.25 | Okay
45 | LMD-PRD | BEAD MILL | CCD-LMD-PR-EQT-06-PS1 | 401 | CHINA | 0.01 | LMD-PRD | 2021-12-24 | 2022-12-24 | 0.05 MPA | 0.6 | 0.65 | -0.05 | Okay
[11 empty] | 0.5 | 0.65 | -0.15 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
46 | LMD-PRD | CHILLER | CCD-LMD-PR-EQT-16-TM1 | 91 | CHINA | 0.01 | LMD-PRD | 2021-12-24 | 2022-12-24 | 5 C | 10 | 8 | 2 | Okay
[11 empty] | 15 | 18 | -3 | Okay
[11 empty] | 15 | 18 | -3 | Okay
47 | LMD-PRD | WEIGHING SCALE | CCD-LMD-PR-EQT-13-WS | 87-A | CHINA | 100 | LMD-PRD | 2021-12-24 | 2022-12-24 | 10G | 9.7 | 10 | -0.3 | Okay
[11 empty] | 9.8 | 10 | -0.2 | Okay
[11 empty] | 9.8 | 10 | -0.2 | Okay
48 | LMD-PRD | WEIGHING SCALE | CCD-LMD-PR-EQT-11-WS | 90 | CHINA | 10 | LMD-PRD | 2021-12-24 | 2022-12-24 | 10G | 9.7 | 10 | -0.3 | Okay
[11 empty] | 9.8 | 10 | -0.2 | Okay
[11 empty] | 9.8 | 10 | -0.2 | Okay
49 | LMD-PRD | MIXER | CCD-LMD-PR-EQT-09-RPM1 | 360 | CHINA | 10 | LMD-PRD | 2021-12-24 | 2022-12-24 | 5RPM | 98 | 100 | -2 | Okay
[11 empty] | 95 | 100 | -5 | Okay
[11 empty] | 101 | 100 | 1 | Okay
50 | LMD-PRD | MIXER | CCD-LMD-PR-EQT-09-PG | 262 | CHINA | 0.01 | LMD-LAB | 2021-12-24 | 2022-12-24 | 0.05 MPA | 0.6 | 0.65 | -0.05 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.4 | 0.65 | -0.25 | Okay
51 | LMD-PRD | MIXER | CCD-LMD-PR-EQT-10-RPM M1 | 361-A | CHINA | 10 | LMD-PRD | 2021-12-24 | 2022-12-24 | 5RPM | 100 | 100 | 0 | Okay
[11 empty] | 101 | 100 | 1 | Okay
[11 empty] | 101 | 100 | 1 | Okay
52 | LMD-PRD | WALL MIXER | CCD-LMD-PR-EQT-08-RPM1 | 263 | CHINA | 10 | LMD-PRD | 2021-12-24 | 2022-12-24 | 5RPM | 95 | 100 | -5 | Okay
[11 empty] | 99 | 100 | -1 | Okay
[11 empty] | 101 | 100 | 1 | Okay
53 | LMD-PRD | DPM MACHINE | CCD-LMD-PR-EQT-19-TIMER 1 | 402 | CHINA | 0.5Sec | LMD-PRD | 2021-12-24 | 2022-12-24 | 0.5Sec | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.6 | 0.5 | 0.1 | Okay
[11 empty] | 0.4 | 0.5 | -0.1 | Okay
54 | LMD-PRD | WEIGHING SCALE | CCD-LMD-PR-EQT-14-WS | 151-A | CHINA | 10 | LMD-PRD | 2021-12-24 | 2022-12-24 | 10G | 9.5 | 10 | -0.5 | Okay
[11 empty] | 9.6 | 10 | -0.4 | Okay
[11 empty] | 9.8 | 10 | -0.2 | Okay
55 | LMD-PRD | MAIN CONTROL PANNEL | CCD-LMD-PR-EQT-22-DIG AMP M1 | 235 | CHINA | 1 | LMD-PRD | 2021-12-24 | 2022-12-24 | 0.5A | 0.4 | 0.5 | -0.1 | Okay
[11 empty] | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
56 | LMD-PRD | MAIN CONTROL PANNEL | CCD-LMD-PR-EQT-22-DIG V M 1 | 234 | CHINA | 1 | LMD-PRD | 2021-12-24 | 2022-12-24 | 0.5V | 0.8 | 0.5 | 0.3 | Okay
[11 empty] | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
57 | LMD-PRD | DPM MACHINE100LTR | CCD-LMD-PR-EQT-23-TIMER 1 | 453 | CHINA | 0.5Sec | LMD-PRD | 2021-12-24 | 2022-12-24 | 0.5Sec | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.6 | 0.5 | 0.1 | Okay
[11 empty] | 0.4 | 0.5 | -0.1 | Okay
58 | LMD-PRD | DPM MACHINE 600LTR | CCD-LMD-PR-EQT-28-TIMER 1 | 454 | CHINA | 0.5Sec | LMD-PRD | 2021-12-24 | 2022-12-24 | 0.5Sec | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.6 | 0.5 | 0.1 | Okay
[11 empty] | 0.4 | 0.5 | -0.1 | Okay
59 | LMD-PRD | HYDRULIC MIXER (ELASTO Sole) | CCD-LMD-PR-EQT-31-RPM1 | 470 | CHINA | 10 | LMD-PRD | 2021-12-24 | 2022-12-24 | 5RPM | 98 | 100 | -2 | Okay
[11 empty] | 95 | 100 | -5 | Okay
[11 empty] | 101 | 100 | 1 | Okay
60 | LMD-PRD | HYDRULIC MIXER (ELASTO Sole) | CCD-LMD-PR-EQT-31-AMP1 | 469 | CHINA | 1 | LMD-PRD | 2021-12-24 | 2022-12-24 | 1A | 0.2 | 0.5 | -0.3 | Okay
[11 empty] | 0.9 | 0.5 | 0.4 | Okay
[11 empty] | 0.1 | 0.5 | -0.4 | Okay
61 | LMD-PRD | WEIGHING SCALE | CCD-LMD-PR-EQT-12-WS | 90-A | CHINA | 10 | LMD-PRD | 2021-12-24 | 2022-12-24 | 10G | 9.7 | 10 | -0.3 | Okay
[11 empty] | 9.8 | 10 | -0.2 | Okay
[11 empty] | 9.8 | 10 | -0.2 | Okay
62 | LMD-PRD | CHILLER | CCD-LMD-PR-EQT-32-TC1 | 494 | CHINA | 0.01 | LMD-PRD | 2021-12-24 | 2022-12-24 | 5 C | 10 | 8 | 2 | Okay
[11 empty] | 15 | 18 | -3 | Okay
[11 empty] | 15 | 18 | -3 | Okay
63 | LMD-LAB | OVEN | CCD-LMD-PR-EQT-33-TC | 495 | CHINA | 0.01 | LMD-PRD | 2021-12-24 | 2022-12-24 | 5 C | 10 | 8 | 2 | Okay
[11 empty] | 15 | 18 | -3 | Okay
[11 empty] | 15 | 18 | -3 | Okay
64 | LMD-LAB | HEAT DYING MACHINE | CCD-LMD-PR-EQT-34-TC | 473 | CHINA | 1 | LMD-PRD | 2021-12-24 | 2022-12-24 | 5 C | 10 | 8 | 2 | Okay
[11 empty] | 15 | 18 | -3 | Okay
[11 empty] | 15 | 18 | -3 | Okay
65 | LMD-LAB | HEAT DYING MACHINE | CCD-LMD-PR-EQT-34-TIMER | 474 | CHINA | 0.5Sec | LMD-LAB | 2021-12-24 | 2022-12-24 | 0.5Sec | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.6 | 0.5 | 0.1 | Okay
[11 empty] | 0.4 | 0.5 | -0.1 | Okay
66 | LMD-LAB | BEAD MILL | CCD-LMD-PR-EQT-35-AMP | 213 | CHINA | 1 | LMD-PRD | 2021-12-24 | 2022-12-24 | 1A | 0.2 | 0.5 | -0.3 | Okay
[11 empty] | 0.9 | 0.5 | 0.4 | Okay
[11 empty] | 0.1 | 0.5 | -0.4 | Okay
67 | LMD-LAB | BEAD MILL | CCD-LMD-PR-EQT-35-TIMER | 214 | CHINA | 0.5Sec | LMD-LAB | 2021-12-24 | 2022-12-24 | 0.5Sec | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.6 | 0.5 | 0.1 | Okay
[11 empty] | 0.4 | 0.5 | -0.1 | Okay
68 | LMD-LAB | BEAD MILL | CCD-LMD-PR-EQT-35-TG | 215 | CHINA | 1 | LMD-PRD | 2021-12-24 | 2022-12-24 | 5 C | 25 | 25 | 0 | Okay
[11 empty] | 45 | 45.1 | -0.1 | Okay
[11 empty] | 60 | 60.3 | -0.3 | Okay
69 | LMD-LAB | BEAD MILL | CCD-LMD-PR-EQT-35-PG | 216 | CHINA | 0.01 | LMD-LAB | 2021-12-24 | 2022-12-24 | 0.05 MPA | 0.6 | 0.65 | -0.05 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.4 | 0.65 | -0.25 | Okay
70 | LMD-LAB | WATER BATH | CCD-LMD-PR-EQT-36-TC | 231 | CHINA | 0.01 | LMD-PRD | 2021-12-24 | 2022-12-24 | 5 C | 10 | 8 | 2 | Okay
[11 empty] | 20 | 18 | 2 | Okay
[11 empty] | 15 | 18 | -3 | Okay
71 | LMD-LAB | PNEUMATIC MIXER | CCD-LMD-LAB-EQT-08-AMP-1 | 236 | CHINA | 0.5A | LMD-LAB | 2021-12-24 | 2022-12-24 | 1A | 0.2 | 0.5 | -0.3 | Okay
[11 empty] | 0.9 | 0.5 | 0.4 | Okay
[11 empty] | 0.1 | 0.5 | -0.4 | Okay
72 | LMD-LAB | PNEUMATIC MIXER | CCD-LMD-LAB-EQT-08-RPM1 | 229 | CHINA | 10 | LMD-LAB | 2021-12-24 | 2022-12-24 | 5RPM | 98 | 100 | -2 | Okay
[11 empty] | 95 | 100 | -5 | Okay
[11 empty] | 101 | 100 | 1 | Okay
73 | LMD-LAB | BEAD MILL | CCD-LMD-LAB-EQT-02-AMP1- | 228 | CHINA | 1 | LMD-LAB | 2021-12-24 | 2022-12-24 | 1A | 0.2 | 0.5 | -0.3 | Okay
[11 empty] | 0.9 | 0.5 | 0.4 | Okay
[11 empty] | 0.1 | 0.5 | -0.4 | Okay
74 | LMD-LAB | BEAD MILL | CCD-LMD-LAB-EQT-02-RPM1 | 227 | CHINA | 10 | LMD-LAB | 2021-12-24 | 2022-12-24 | 5RPM | 101 | 100 | 1 | Okay
[11 empty] | 100 | 100 | 0 | Okay
[11 empty] | 101 | 100 | 1 | Okay
75 | LMD-LAB | BEAD MILL | CCD-LMD-LAB-EQT-01-AMP1 | 226 | CHINA | 1 | LMD-LAB | 2021-12-24 | 2022-12-24 | 1A | 0.2 | 0.5 | -0.3 | Okay
[11 empty] | 0.9 | 0.5 | 0.4 | Okay
[11 empty] | 0.1 | 0.5 | -0.4 | Okay
76 | LMD-LAB | BEAD MILL | CCD-LMD-LAB-EQT-01-RPM1 | 225 | CHINA | 10 | LMD-LAB | 2021-12-24 | 2022-12-24 | 5RPM | 101 | 100 | 1 | Okay
[11 empty] | 100 | 100 | 0 | Okay
[11 empty] | 101 | 100 | 1 | Okay
77 | LMD-LAB | PNEUMATIC MIXER | CCD-LMD-LAB-EQT-07-AMP1 | 224 | CHINA | 1 | LMD-LAB | 2021-12-24 | 2022-12-24 | 1A | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.8 | 0.5 | 0.3 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
78 | LMD-LAB | PNEUMATIC MIXER | CCD-LMD-LAB-EQT-07-RPM1 | 223 | CHINA | 10 | LMD-LAB | 2021-12-24 | 2022-12-24 | 10RPM | 101 | 100 | 1 | Okay
[11 empty] | 100 | 100 | 0 | Okay
[11 empty] | 101 | 100 | 1 | Okay
79 | LMD-LAB | BEAD MILL | CCD-LMD-LAB-EQT-03-PG1 | 219 | CHINA | 0.01 | LMD-LAB | 2021-12-24 | 2022-12-24 | 0.05 MPA | 0.6 | 0.65 | -0.05 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.4 | 0.65 | -0.25 | Okay
80 | LMD-LAB | BEAD MILL | CCD-LMD-LAB-EQT-03-PG2 | 221 | CHINA | 0.01 | LMD-LAB | 2021-12-24 | 2022-12-24 | 0.05 MPA | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.4 | 0.65 | -0.25 | Okay
[11 empty] | 0.6 | 0.65 | -0.05 | Okay
81 | LMD-LAB | BEAD MILL | CCD-LMD-LAB-EQT-03-PG3 | 222 | CHINA | 0.01 | LMD-LAB | 2021-12-24 | 2022-12-24 | 0.05 MPA | 0.6 | 0.65 | -0.05 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.4 | 0.65 | -0.25 | Okay
82 | LMD-LAB | BEAD MILL | CCD-LMD-LAB-EQT-03-AMP | 217 | CHINA | 0.5A | LMD-LAB | 2021-12-24 | 2022-12-24 | 1A | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.8 | 0.5 | 0.3 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
83 | LMD-LAB | BEAD MILL | CCD-LMD-LAB-EQT-03-TG1 | 220 | CHINA | 1 | LMD-LAB | 2021-12-24 | 2022-12-24 | 5 C | 15 | 15 | 0 | Okay
[11 empty] | 35 | 35.2 | -0.2 | Okay
[11 empty] | 50 | 50.4 | -0.4 | Okay
84 | LMD-LAB | BEAD MILL | CCD-LMD-LAB-EQT-03-TIMER | 218 | CHINA | 0.5Sec | LMD-LAB | 2021-12-24 | 2022-12-24 | 0.5Sec | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.6 | 0.5 | 0.1 | Okay
[11 empty] | 0.4 | 0.5 | -0.1 | Okay
85 | LMD-LAB | DPM | CCD-LMD-LAB-EQT-17-TIMER | 451 | CHINA | 0.5Sec | LMD-LAB | 2021-12-24 | 2022-12-24 | 0.5Sec | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.6 | 0.5 | 0.1 | Okay
[11 empty] | 0.4 | 0.5 | -0.1 | Okay
86 | LMD-LAB | DPM | CCD-LMD-LAB-EQT-18-TIMER | 452 | CHINA | 0.5Sec | LMD-LAB | 2021-12-24 | 2022-12-24 | 0.5Sec | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.6 | 0.5 | 0.1 | Okay
[11 empty] | 0.4 | 0.5 | -0.1 | Okay
87 | LMD-LAB | HIGH SPEED DISPERSER 0.37 KW | CCD-LMD-LAB-EQT--26-AMP | 465 | CHINA | 0.5A | LMD-LAB | 2021-12-24 | 2022-12-24 | 1A | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.8 | 0.5 | 0.3 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
88 | LMD-LAB | HIGH SPEED DISPERSER 0.37 KW | CCD-LMD-LAB-EQT-26-RPM | 466 | CHINA | 10 | LMD-LAB | 2021-12-24 | 2022-12-24 | 5RPM | 101 | 100 | 1 | Okay
[11 empty] | 100 | 100 | 0 | Okay
[11 empty] | 101 | 100 | 1 | Okay
89 | LMD-LAB | HIGH SPEED DISPERSER 0.37 KW | CCD-LMD-LAB-EQT-27-RPM | 468 | CHINA | 10 | LMD-LAB | 2021-12-24 | 2022-12-24 | 5RPM | 98 | 100 | -2 | Okay
[11 empty] | 95 | 100 | -5 | Okay
[11 empty] | 101 | 100 | 1 | Okay
90 | LMD-LAB | HIGH SPEED DISPERSER 0.37 KW | CCD-LMD-LAB-EQT-27-AMP | 467 | CHINA | 0.5A | LMD-LAB | 2021-12-24 | 2022-12-24 | 1A | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.8 | 0.5 | 0.3 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
91 | LMD-LAB | TWO ROLL MILL | CCD-LMD-LAB-EQT-04-TEM M | 471 | CHINA | 0.01 | LMD-LAB | 2021-12-24 | 2022-12-24 | 5 C | 10 | 8 | 2 | Okay
[11 empty] | 14 | 18 | -4 | Okay
[11 empty] | 15 | 18 | -3 | Okay
92 | LMD-LAB | TWO ROLL MILL | CCD-LMD-LAB-EQT-04-TEM M | 472 | CHINA | 0.01 | LMD-LAB | 2021-12-24 | 2022-12-24 | 5 C | 10 | 8 | 2 | Okay
[11 empty] | 15 | 18 | -3 | Okay
[11 empty] | 14 | 18 | -4 | Okay
93 | LMD-LAB | TWO ROLL MILL | CCD-LMD-LAB-EQT-04-RPM | 473 | CHINA | 10 | LMD-LAB | 2021-12-24 | 2022-12-24 | 5RPM | 101 | 100 | 1 | Okay
[11 empty] | 100 | 100 | 0 | Okay
[11 empty] | 101 | 100 | 1 | Okay
94 | LMD-LAB | BEAD MILL | CCD-LMD-LAB-EQT-13-HRM | 211 | CHINA | 0.5Sec | LMD-LAB | 2021-12-24 | 2022-12-24 | 0.5Sec | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.6 | 0.5 | 0.1 | Okay
[11 empty] | 0.4 | 0.5 | -0.1 | Okay
95 | LMD-LAB | BEAD MILL | CCD-LMD-LAB-EQT-13-A1 | 210 | CHINA | 1 | LMD-LAB | 2021-12-24 | 2022-12-24 | 1A | 0.4 | 0.5 | -0.1 | Okay
[11 empty] | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
96 | LMD-LAB | BEAD MILL | CCD-LMD-LAB-EQT-13-PS | 209 | CHINA | 1 | LMD-LAB | 2021-12-24 | 2022-12-24 | 5 C | 15 | 15 | 0 | Okay
[11 empty] | 35 | 35.2 | -0.2 | Okay
[11 empty] | 50 | 50.4 | -0.4 | Okay
97 | LMD-LAB | BEAD MILL | CCD-LMD-LAB-EQT-13-PG1 | 208 | CHINA | 0.01 | LMD-LAB | 2021-12-24 | 2022-12-24 | 0.05 MPA | 0.6 | 0.65 | -0.05 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.4 | 0.65 | -0.25 | Okay
98 | LMD-LAB | BEAD MILL | CCD-LMD-LAB-EQT-13-PG2 | 207 | CHINA | 0.01 | LMD-LAB | 2021-12-24 | 2022-12-24 | 0.05 MPA | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.4 | 0.65 | -0.25 | Okay
[11 empty] | 0.6 | 0.65 | -0.05 | Okay
99 | LMD-LAB | WEIGHING SCALE | CCD-LMD-LAB-EQT-11-WS | 386 | CHINA | 0.01GM | LMD-LAB | 2021-12-24 | 2022-12-24 | 10G | 9.5 | 10 | -0.5 | Okay
[11 empty] | 9.6 | 10 | -0.4 | Okay
[11 empty] | 9.8 | 10 | -0.2 | Okay
100 | LMD-LAB | WEIGHING SCALE | CCD-LMD-LAB-EQT-12-WS | 387 | CHINA | 0.01GM | LMD-LAB | 2021-12-24 | 2022-12-24 | 10G | 9.5 | 10 | -0.5 | Okay
[11 empty] | 9.6 | 10 | -0.4 | Okay
[11 empty] | 9.8 | 10 | -0.2 | Okay
101 | LMD-LAB | OVEN | CCD-LMD-LAB-EQT-15-OVEN | 71 | CHINA | 1 | LMD-LAB | 2021-12-24 | 2022-12-24 | 5 C | 15 | 15 | 0 | Okay
[11 empty] | 35 | 35.2 | -0.2 | Okay
[11 empty] | 50 | 50.4 | -0.4 | Okay
102 | LMD-LAB | SIMPLE MIXER | CCD-LMD-LAB-EQT-09-RPM-01 | 204 | CHINA | 10 | LMD-LAB | 2021-12-24 | 2022-12-24 | 5RPM | 101 | 100 | 1 | Okay
[11 empty] | 100 | 100 | 0 | Okay
[11 empty] | 101 | 100 | 1 | Okay
103 | LMD-LAB | SIMPLE MIXER | CCD-LMD-LAB-EQT-09-AMP1 | 388 | CHINA | 1 | LMD-LAB | 2021-12-24 | 2022-12-24 | 1A | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.8 | 0.5 | 0.3 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
104 | LMD-LAB | SIMPLE MIXER | CCD-LMD-LAB-EQT-10-RPM1 | 206 | CHINA | 10 | LMD-LAB | 2021-12-24 | 2022-12-24 | 5RPM | 99 | 100 | -1 | Okay
[11 empty] | 100 | 100 | 0 | Okay
[11 empty] | 98 | 100 | -2 | Okay
105 | LMD-LAB | SIMPLE MIXER | CCD-LMD-LAB-EQT-10-AMP1 | 205 | CHINA | 1 | LMD-LAB | 2021-12-24 | 2022-12-24 | 1A | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.8 | 0.5 | 0.3 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
106 | LMD-LAB | SUPPLY BOX | CCD-LMD-LAB-EQT-24A-DIG V M 1 | 232 | CHINA | 1 | LMD-LAB | 2021-12-24 | 2022-12-24 | 0.5V | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.8 | 0.5 | 0.3 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
107 | LMD-LAB | SUPPLY BOX | CCD-LMD-LAB-EQT-24B-DIG V M 1 | 233 | CHINA | 1 | LMD-LAB | 2021-12-24 | 2022-12-24 | 0.5V | 0.6 | 0.5 | 0.1 | Okay
[11 empty] | 0.4 | 0.5 | -0.1 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
108 | ADMIN HALL | ELECTRICAL SUPPLY BOX | ADMIN HALL EQT-02-DIG V M 1 | 408 | CHINA | 1 | ADMIN HALL | 2021-12-24 | 2022-12-24 | 1A | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.8 | 0.5 | 0.3 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
109 | ADMIN HALL | ELECTRICAL SUPPLY BOX | ADMIN HALL EQT-01-DIG V M 1 | 408-A | CHINA | 1 | ADMIN HALL | 2021-12-24 | 2022-12-24 | 1A | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.8 | 0.5 | 0.3 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
110 | STORE | ELECTRICAL SUPPLY BOX | STORE-U2-EQT-7-DIG V M 1 | 267-B | CHINA | 1 | LMD-LAB | 2021-12-24 | 2022-12-24 | 1A | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.8 | 0.5 | 0.3 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
111 | STORE | WEIGHING SCALE | STORE-U2-EQT-01-WS | 88 | CHINA | 10 | STORE | 2021-12-24 | 2022-12-24 | 10G | 9.5 | 10 | -0.5 | Okay
[11 empty] | 9.6 | 10 | -0.4 | Okay
[11 empty] | 9.8 | 10 | -0.2 | Okay
112 | STORE | WEIGHING SCALE | STORE-U2-EQT-02-WS | 89 | CHINA | 10 | STORE | 2021-12-24 | 2022-12-24 | 10G | 9.7 | 10 | -0.3 | Okay
[11 empty] | 9.6 | 10 | -0.4 | Okay
[11 empty] | 9.8 | 10 | -0.2 | Okay
113 | STORE | WEIGHING SCALE | STORE-U2-EQT-03-WS | 87-B | CHINA | 100 | STORE | 2021-12-24 | 2022-12-24 | 10G | 9.2 | 10 | -0.8 | Okay
[11 empty] | 9.6 | 10 | -0.4 | Okay
[11 empty] | 9.8 | 10 | -0.2 | Okay
75 | LMD-LAB | BEAD MILL | CCD-LMD-LAB-EQT-40-AMP1 | 301 | CHINA | 1 | LMD-LAB | 2021-12-24 | 2022-12-24 | 1A | 0.2 | 0.5 | -0.3 | Okay
[11 empty] | 0.9 | 0.5 | 0.4 | Okay
[11 empty] | 0.1 | 0.5 | -0.4 | Okay
76 | LMD-LAB | BEAD MILL | CCD-LMD-LAB-EQT-40-RPM1 | 302 | CHINA | 10 | LMD-LAB | 2021-12-24 | 2022-12-24 | 5RPM | 101 | 100 | 1 | Okay
[11 empty] | 100 | 100 | 0 | Okay
[11 empty] | 101 | 100 | 1 | Okay
75 | LMD-LAB | BEAD MILL | CCD-LMD-LAB-EQT-41-AMP1 | 303 | CHINA | 1 | LMD-LAB | 2021-12-24 | 2022-12-24 | 1A | 0.2 | 0.5 | -0.3 | Okay
[11 empty] | 0.9 | 0.5 | 0.4 | Okay
[11 empty] | 0.1 | 0.5 | -0.4 | Okay
76 | LMD-LAB | BEAD MILL | CCD-LMD-LAB-EQT-41-RPM1 | 304 | CHINA | 10 | LMD-LAB | 2021-12-24 | 2022-12-24 | 5RPM | 101 | 100 | 1 | Okay
[11 empty] | 100 | 100 | 0 | Okay
[11 empty] | 101 | 100 | 1 | Okay
75 | LMD-LAB | BEAD MILL | CCD-LMD-LAB-EQT-42-AMP1 | 305 | CHINA | 1 | LMD-LAB | 2021-12-24 | 2022-12-24 | 1A | 0.2 | 0.5 | -0.3 | Okay
[11 empty] | 0.9 | 0.5 | 0.4 | Okay
[11 empty] | 0.1 | 0.5 | -0.4 | Okay
76 | LMD-LAB | BEAD MILL | CCD-LMD-LAB-EQT-42-RPM1 | 306 | CHINA | 10 | LMD-LAB | 2021-12-24 | 2022-12-24 | 5RPM | 101 | 100 | 1 | Okay
[11 empty] | 100 | 100 | 0 | Okay
[11 empty] | 101 | 100 | 1 | Okay
49 | LMD-PRD | Hydraulic MIXER 4 | CCD-LMD-PR-EQT-29-RPM1 | 307 | CHINA | 10 | LMD-PRD | 2021-12-24 | 2022-12-24 | 5RPM | 98 | 100 | -2 | Okay
[11 empty] | 95 | 100 | -5 | Okay
[11 empty] | 101 | 100 | 1 | Okay
50 | LMD-PRD | Hydraulic MIXER 4 | CCD-LMD-PR-EQT-29-PG | 308 | CHINA | 0.01 | LMD-LAB | 2021-12-24 | 2022-12-24 | 0.05 MPA | 0.6 | 0.65 | -0.05 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.4 | 0.65 | -0.25 | Okay
[8 empty] | 2021-12-24 | 2022-12-24
49 | LMD-PRD | Hydraulic MIXER 5 | CCD-LMD-PR-EQT-30-RPM1 | 309 | CHINA | 10 | LMD-PRD | 2021-12-24 | 2022-12-24 | 5RPM | 98 | 100 | -2 | Okay
[11 empty] | 95 | 100 | -5 | Okay
[11 empty] | 101 | 100 | 1 | Okay
50 | LMD-PRD | Hydraulic MIXER 5 | CCD-LMD-PR-EQT-30-PG | 310 | CHINA | 0.01 | LMD-LAB | 2021-12-24 | 2022-12-24 | 0.05 MPA | 0.6 | 0.65 | -0.05 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.4 | 0.65 | -0.25 | Okay
49 | LMD-PRD | DPM MIXER | CCD-LMD-PR-EQT-33-RPM1 | 311 | CHINA | 10 | LMD-PRD | 2021-12-24 | 2022-12-24 | 5RPM | 98 | 100 | -2 | Okay
[11 empty] | 95 | 100 | -5 | Okay
[11 empty] | 101 | 100 | 1 | Okay
50 | LMD-PRD | DPM MIXER | CCD-LMD-PR-EQT-33-TS | 312 | CHINA | 0.01 | LMD-LAB | 2021-12-24 | 2022-12-24 | 0.05 MPA | 0.6 | 0.65 | -0.05 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.4 | 0.65 | -0.25 | Okay
Calibrated By: | [11 empty] | Verified By:
Abrivations
1 | TG |  | TEMPRATURE GAUGE
2 | PG |  | PRESSURE GAUGE
3 | AMP M |  | AMPERE METER
4 | TIMER |  | TIMER
5 | TM |  | TEMPRATURE METER
6 | TS |  | TEMPRATUER SENSOR
7 | WS |  | WEIGHING SCALE
8 | RPM M |  | REVOLOUTION PER MINUTE METER
9 | DIG AMP M |  | DIGITAL AMPERE METER
10 | VM |  | VOLT METER
11 | A |  | AMPERE METER
12 | OVEN |  | OVEN
13 | DIG V M |  | DIGITAL VOLT METER
14 | TEMP BEN M |  | TEMPRATURE BENCH METER
15 | LIG B |  | LIGHT BOX
16 | HSV |  | HYDRULIC SOLINOID VALVE
17 | RFM

## PCC
DOC #: BRCC/MGT/FRM-053-U2 ISSUE STATUS: 02 ISSUE DATE: Jan 24, 2019
In House Calibration Records
Master Standard Value
Sr. | Department | MACHINE NAME | Instrument Name | Calibration Code | Make / Model | Range/Least Count | Location | Calibration Date | Next Calibration Due Date | Deviation limit | Calibration Results |  |  | Remarks
[11 empty] | Standard Value | Instrument under calibration | Deviation
261 | PCC-PRD | BEAD MILL | PCC-PR-EQT-01-PG1 | 179 | CHINA | 0.01 | PCC-PRD | 2021-12-23 | 2022-12-23 | 0.05 MPA | 0.5 | 0.65 | -0.15 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.4 | 0.65 | -0.25 | Okay
262 | PCC-PRD | BEAD MILL | PCC-PR-EQT-01-PG2 | 178 | CHINA | 0.01 | PCC-PRD | 2021-12-23 | 2022-12-23 | 0.05 MPA | 0.6 | 0.65 | -0.05 | Okay
[11 empty] | 0.4 | 0.65 | -0.25 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
263 | PCC-PRD | BEAD MILL | PCC-PR-EQT-01-PG3 | 92 | CHINA | 0.01 | PCC-PRD | 2021-12-23 | 2022-12-23 | 0.05 MPA | 0.5 | 0.65 | -0.15 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.4 | 0.65 | -0.25 | Okay
264 | PCC-PRD | BEAD MILL | PCC-PR-EQT-01-PG4 | 385 | CHINA | 0.01 | PCC-PRD | 2021-12-23 | 2022-12-23 | 0.05 MPA | 0.6 | 0.65 | -0.05 | Okay
[11 empty] | 0.4 | 0.65 | -0.25 | Okay
[11 empty] | 0.63 | 0.65 | -0.02 | Okay
265 | PCC-PRD | BEAD MILL | PCC-PR-EQT-01-TG1 | 97 | CHINA | 1 | PCC-PRD | 2021-12-23 | 2022-12-23 | 5 C | 25 | 25 | 0 | Okay
[11 empty] | 45 | 45.1 | -0.1 | Okay
[11 empty] | 60 | 60.3 | -0.3 | Okay
266 | PCC-PRD | BEAD MILL | PCC-PR-EQT-01-AMP 1 | 177 | CHINA | 1 | PCC-PRD | 2021-12-23 | 2022-12-23 | 1A | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
[11 empty] | 0.8 | 0.5 | 0.3 | Okay
267 | PCC-PRD | BEAD MILL | PCC-PR-EQT-07-AMP1 | 180 | CHINA | 1 | PCC-PRD | 2021-12-23 | 2022-12-23 | 1A | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
[11 empty] | 0.8 | 0.5 | 0.3 | Okay
268 | PCC-PRD | BEAD MILL | PCC-PR-EQT-07-PG1 | 181 | CHINA | 0.01 | PCC-PRD | 2021-12-23 | 2022-12-23 | 0.05 MPA | 0.5 | 0.65 | -0.15 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.4 | 0.65 | -0.25 | Okay
269 | PCC-PRD | BEAD MILL | PCC-PR-EQT-07-PG2 | 93 | CHINA | 0.01 | PCC-PRD | 2021-12-23 | 2022-12-23 | 0.05 MPA | 0.5 | 0.65 | -0.15 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.4 | 0.65 | -0.25 | Okay
270 | PCC-PRD | BEAD MILL | PCC-PR-EQT-07-PG3 | 182 | CHINA | 0.01 | PCC-PRD | 2021-12-23 | 2022-12-23 | 0.05 MPA | 0.5 | 0.65 | -0.15 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.4 | 0.65 | -0.25 | Okay
271 | PCC-PRD | BEAD MILL | PCC-PR-EQT-07-PG4 | 384 | CHINA | 0.01 | PCC-PRD | 2021-12-23 | 2022-12-23 | 0.05 MPA | 0.5 | 0.65 | -0.15 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.4 | 0.65 | -0.25 | Okay
272 | PCC-PRD | BEAD MILL | PCC-PR-EQT-07-TG1 | 98 | CHINA | 1 | PCC-PRD | 2021-12-23 | 2022-12-23 | 5 C | 25 | 25 | 0 | Okay
[11 empty] | 45 | 45.1 | -0.1 | Okay
[11 empty] | 60 | 60.3 | -0.3 | Okay
273 | PCC-PRD | BEAD MILL | PCC-PR-EQT-12-AMP1 | 183 | CHINA | 1 | PCC-PRD | 2021-12-23 | 2022-12-23 | 1A | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
[11 empty] | 0.8 | 0.5 | 0.3 | Okay
274 | PCC-PRD | BEAD MILL | PCC-PR-EQT-12-PG1 | 94 | CHINA | 0.01 | PCC-PRD | 2021-12-23 | 2022-12-23 | 0.05 MPA | 0.5 | 0.65 | -0.15 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.4 | 0.65 | -0.25 | Okay
275 | PCC-PRD | BEAD MILL | PCC-PR-EQT-12-PG2 | 184 | CHINA | 0.01 | PCC-PRD | 2021-12-23 | 2022-12-23 | 0.05 MPA | 0.6 | 0.65 | -0.05 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.5 | 0.65 | -0.15 | Okay
276 | PCC-PRD | BEAD MILL | PCC-PR-EQT-12-PG3 | 185 | CHINA | 0.01 | PCC-PRD | 2021-12-23 | 2022-12-23 | 0.05 MPA | 0.5 | 0.65 | -0.15 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.4 | 0.65 | -0.25 | Okay
277 | PCC-PRD | BEAD MILL | PCC-PR-EQT-12-PG4 | 383 | CHINA | 0.01 | PCC-PRD | 2021-12-23 | 2022-12-23 | 0.05 MPA | 0.5 | 0.65 | -0.15 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.4 | 0.65 | -0.25 | Okay
278 | PCC-PRD | BEAD MILL | PCC-PR-EQT-12-TG1 | 99 | CHINA | 1 | PCC-PRD | 2021-12-23 | 2022-12-23 | 5 C | 25 | 25 | 0 | Okay
[11 empty] | 45 | 45.1 | -0.1 | Okay
[11 empty] | 60 | 60.3 | -0.3 | Okay
279 | PCC-PRD | BEAD MILL | PCC-PR-EQT-13-AMP1 | 186 | CHINA | 1 | PCC-PRD | 2021-12-23 | 2022-12-23 | 1A | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
[11 empty] | 0.8 | 0.5 | 0.3 | Okay
280 | PCC-PRD | BEAD MILL | PCC-PR-EQT-13-PG1 | 187 | CHINA | 0.01 | PCC-PRD | 2021-12-23 | 2022-12-23 | 0.05 MPA | 0.4 | 0.65 | -0.25 | Okay
[11 empty] | 0.6 | 0.65 | -0.05 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
281 | PCC-PRD | BEAD MILL | PCC-PR-EQT-13-PG2 | 188 | CHINA | 0.01 | PCC-PRD | 2021-12-23 | 2022-12-23 | 0.05 MPA | 0.6 | 0.65 | -0.05 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.4 | 0.65 | -0.25 | Okay
282 | PCC-PRD | BEAD MILL | PCC-PR-EQT-13-PG3 | 95 | CHINA | 0.01 | PCC-PRD | 2021-12-23 | 2022-12-23 | 0.05 MPA | 0.4 | 0.65 | -0.25 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.6 | 0.65 | -0.05 | Okay
283 | PCC-PRD | BEAD MILL | PCC-PR-EQT-13-PG4 | 382 | CHINA | 0.01 | PCC-PRD | 2021-12-23 | 2022-12-23 | 0.05 MPA | 0.4 | 0.65 | -0.25 | Okay
[11 empty] | 0.6 | 0.65 | -0.05 | Okay
[11 empty] | 0.4 | 0.65 | -0.25 | Okay
284 | PCC-PRD | BEAD MILL | PCC-PR-EQT-13-TG1 | 100 | CHINA | 1 | PCC-PRD | 2021-12-23 | 2022-12-23 | 5 C | 20 | 20 | 0 | Okay
[11 empty] | 40 | 40.3 | -0.3 | Okay
[11 empty] | 55 | 55.1 | -0.1 | Okay
285 | PCC-PRD | BEAD MILL | PCC-PR-EQT-14-AMP1 | 191 | CHINA | 1 | PCC-PRD | 2021-12-23 | 2022-12-23 | 1A | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
[11 empty] | 0.8 | 0.5 | 0.3 | Okay
286 | PCC-PRD | BEAD MILL | PCC-PR-EQT-14-PG1 | 96 | CHINA | 0.01 | PCC-PRD | 2021-12-23 | 2022-12-23 | 0.05 MPA | 0.4 | 0.65 | -0.25 | Okay
[11 empty] | 0.6 | 0.65 | -0.05 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
287 | PCC-PRD | BEAD MILL | PCC-PR-EQT-14-PG2 | 189 | CHINA | 0.01 | PCC-PRD | 2021-12-23 | 2022-12-23 | 0.05 MPA | 0.5 | 0.65 | -0.15 | Okay
[11 empty] | 0.5 | 0.65 | -0.15 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
288 | PCC-PRD | BEAD MILL | PCC-PR-EQT-14-PG3 | 154 | CHINA | 0.01 | PCC-PRD | 2021-12-23 | 2022-12-23 | 0.05 MPA | 0.4 | 0.65 | -0.25 | Okay
[11 empty] | 0.5 | 0.65 | -0.15 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
289 | PCC-PRD | BEAD MILL | PCC-PR-EQT-14-PG4 | 381 | CHINA | 0.01 | PCC-PRD | 2021-12-23 | 2022-12-23 | 0.05 MPA | 0.6 | 0.65 | -0.05 | Okay
[11 empty] | 0.5 | 0.65 | -0.15 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
290 | PCC-PRD | BEAD MILL | PCC-PR-EQT-14-TG1 | 101 | CHINA | 1 | PCC-PRD | 2021-12-23 | 2022-12-23 | 5 C | 20 | 20 | 0 | Okay
[11 empty] | 40 | 40.3 | -0.3 | Okay
[11 empty] | 55 | 55.1 | -0.1 | Okay
291 | PCC-PRD | BEAD MILL | PCC-PR-EQT-20-AMP1 | 193-A | CHINA | 1 | PCC-PRD | 2021-12-23 | 2022-12-23 | 1A | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
[11 empty] | 0.8 | 0.5 | 0.3 | Okay
292 | PCC-PRD | BEAD MILL | PCC-PR-EQT-20-PG1 | 74-A | CHINA | 0.01 | PCC-PRD | 2021-12-23 | 2022-12-23 | 0.05 MPA | 0.5 | 0.65 | -0.15 | Okay
[11 empty] | 0.5 | 0.65 | -0.15 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
293 | PCC-PRD | BEAD MILL | PCC-PR-EQT-20-PG2 | 73-A | CHINA | 0.01 | PCC-PRD | 2021-12-23 | 2022-12-23 | 0.05 MPA | 0.4 | 0.65 | -0.25 | Okay
[11 empty] | 0.5 | 0.65 | -0.15 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
294 | PCC-PRD | BEAD MILL | PCC-PR-EQT-20-PG3 | 74-B | CHINA | 0.01 | PCC-PRD | 2021-12-23 | 2022-12-23 | 0.05 MPA | 0.6 | 0.65 | -0.05 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.5 | 0.65 | -0.15 | Okay
295 | PCC-PRD | BEAD MILL | PCC-PR-EQT-20-PS | 194-B | CHINA | ON/Off | PCC-PRD | 2021-12-23 | 2022-12-23 | 0.05 MPA | 0.5 | 0.65 | -0.15 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.6 | 0.65 | -0.05 | Okay
296 | PCC-PRD | BEAD MILL | PCC-PR-EQT-20-TG1 | 193-B | CHINA | 1 | PCC-PRD | 2021-12-23 | 2022-12-23 | 5 C | 20 | 20 | 0 | Okay
[11 empty] | 40 | 40.3 | -0.3 | Okay
[11 empty] | 55 | 55.1 | -0.1 | Okay
297 | PCC-PRD | BEAD MILL | PCC-PR-EQT-20-TG2 | 378-A | CHINA | 1 | PCC-PRD | 2021-12-23 | 2022-12-23 | 5 C | 15 | 15 | 0 | Okay
[11 empty] | 35 | 35.3 | -0.3 | Okay
[11 empty] | 50 | 50.1 | -0.1 | Okay
298 | PCC-PRD | BEAD MILL | PCC-PR-EQT-20-TG3 | 379-A | CHINA | 1 | PCC-PRD | 2021-12-23 | 2022-12-23 | 5 C | 20 | 20 | 0 | Okay
[11 empty] | 40 | 40.3 | -0.3 | Okay
[11 empty] | 55 | 55.1 | -0.1 | Okay
299 | PCC-PRD | BEAD MILL | PCC-PR-EQT-20-TG3 | 194-A | CHINA | 1 | PCC-PRD | 2021-12-23 | 2022-12-23 | 5 C | 20 | 20 | 0 | Okay
[11 empty] | 40 | 40.3 | -0.3 | Okay
[11 empty] | 55 | 55.1 | -0.1 | Okay
300 | PCC-PRD | RO PLANT | PCC-PR-EQT-32-AMP1 | 75 | CHINA | 1 | PCC-PRD | 2021-12-23 | 2022-12-23 | 1A | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
[11 empty] | 0.8 | 0.5 | 0.3 | Okay
301 | PCC-PRD | RO PLANT | PCC-PR-EQT-32-PG1 | 106 | CHINA | 0.01 | PCC-PRD | 2021-12-23 | 2022-12-23 | 0.05 MPA | 0.6 | 0.65 | -0.05 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.4 | 0.65 | -0.25 | Okay
302 | PCC-PRD | RO PLANT | PCC-PR-EQT-32-PG2 | 107 | CHINA | 0.01 | PCC-PRD | 2021-12-23 | 2022-12-23 | 0.05 MPA | 0.4 | 0.65 | -0.25 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.6 | 0.65 | -0.05 | Okay
303 | PCC-PRD | WEIGHING SCALE | PCC-PR-EQT-22-WS | 484 | CHINA | 10 | PCC-PRD | 2021-12-23 | 2022-12-23 | 10G | 9.9 | 10 | -0.1 | Okay
[11 empty] | 10.3 | 10 | 0.3 | Okay
[11 empty] | 9.8 | 10 | -0.2 | Okay
304 | PCC-PRD | WEIGHING SCALE | PCC-PR-EQT-23-WS | 485 | CHINA | 10 | PCC-PRD | 2021-12-23 | 2022-12-23 | 10G | 9.9 | 10 | -0.1 | Okay
[11 empty] | 10.3 | 10 | 0.3 | Okay
[11 empty] | 9.8 | 10 | -0.2 | Okay
305 | PCC-PRD | WEIGHING SCALE | PCC-PR-EQT-10-WS | 102-A | CHINA | 10 | PCC-PRD | 2021-12-23 | 2022-12-23 | 10G | 9.9 | 10 | -0.1 | Okay
[11 empty] | 10.3 | 10 | 0.3 | Okay
[11 empty] | 9.8 | 10 | -0.2 | Okay
306 | PCC-PRD | MIXER | PCC-PR-EQT-16-RPM1 | 195 | CHINA | 10 | PCC-PRD | 2021-12-23 | 2022-12-23 | 5RPM | 101 | 100 | 1 | Okay
[11 empty] | 104 | 100 | 4 | Okay
[11 empty] | 101 | 100 | 1 | Okay
307 | PCC-PRD | MIXER | PCC-PR-EQT-15-RPM1 | 196 | CHINA | 10 | PCC-PRD | 2021-12-23 | 2022-12-23 | 5RPM | 100 | 100 | 0 | Okay
[11 empty] | 98 | 100 | -2 | Okay
[11 empty] | 101 | 100 | 1 | Okay
342 | PCC-PRD | Mixxing Vessel(1) | PCC-PR-EQT-18-RPM | 610 | CHINA | 10 | PCC-LAB | 2021-12-23 | 2022-12-23 | 5RPM | 101 | 100 | 1 | Okay
[11 empty] | 100 | 100 | 0 | Okay
[11 empty] | 101 | 100 | 1 | Okay
343 | PCC-PRD | Mixxing Vessel(1) | PCC-PR-EQT-18-AMP | 611 | CHINA | 0.5A | PCC-LAB | 2021-12-23 | 2022-12-23 | 1A | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.8 | 0.5 | 0.3 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
342 | PCC-PRD | Mixxing Vessel(1) | PCC-PR-EQT-19-RPM | 612 | CHINA | 10 | PCC-LAB | 2021-12-23 | 2022-12-23 | 5RPM | 101 | 100 | 1 | Okay
[11 empty] | 100 | 100 | 0 | Okay
[11 empty] | 101 | 100 | 1 | Okay
343 | PCC-PRD | Mixxing Vessel(1) | PCC-PR-EQT-19-AMP | 613 | CHINA | 0.5A | PCC-LAB | 2021-12-23 | 2022-12-23 | 1A | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.8 | 0.5 | 0.3 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
308 | PCC-PRD | MIXER | PCC-PR-EQT-03-RPM1 | 173 | CHINA | 10 | PCC-PRD | 2021-12-23 | 2022-12-23 | 5RPM | 101 | 100 | 1 | Okay
[11 empty] | 102 | 100 | 2 | Okay
[11 empty] | 101 | 100 | 1 | Okay
309 | PCC-PRD | MIXER | PCC-PR-EQT-03-AMP1 | 174 | CHINA | 1 | PCC-PRD | 2021-12-23 | 2022-12-23 | 1A | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
[11 empty] | 0.8 | 0.5 | 0.3 | Okay
310 | PCC-PRD | MIXER | PCC-PR-EQT-04-AMP1 | 172 | CHINA | 1 | PCC-PRD | 2021-12-23 | 2022-12-23 | 1A | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
[11 empty] | 0.8 | 0.5 | 0.3 | Okay
311 | PCC-PRD | MIXER | PCC-PR-EQT-04-RPM1 | 171 | CHINA | 10 | PCC-PRD | 2021-12-23 | 2022-12-23 | 5RPM | 101 | 100 | 1 | Okay
[11 empty] | 102 | 100 | 2 | Okay
[11 empty] | 101 | 100 | 1 | Okay
312 | PCC-PRD | WEIGHING SCALE | PCC-PR-EQT-06-WS | 108 | CHINA | 10 | PCC-PRD | 2021-12-23 | 2022-12-23 | 10G | 9.9 | 10 | -0.1 | Okay
[11 empty] | 10.6 | 10 | 0.6 | Okay
[11 empty] | 9.8 | 10 | -0.2 | Okay
313 | PCC-PRD | MAIN PANNEL | PCC-PR-EQT-21-VM 1 | 156 | CHINA | 1 | PCC-PRD | 2021-12-23 | 2022-12-23 | 0.5V | 0.4 | 0.5 | -0.1 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
[11 empty] | 0.8 | 0.5 | 0.3 | Okay
314 | PCC-PRD | HYDRULIC MIXER | PCC-PR-EQT-27-RPM | 476 | CHINA | 10 | PCC-PRD | 2021-12-23 | 2022-12-23 | 5RPM | 101 | 100 | 1 | Okay
[11 empty] | 102 | 100 | 2 | Okay
[11 empty] | 101 | 100 | 1 | Okay
315 | PCC-PRD | HYDRULIC MIXER | PCC-PR-EQT-28-RPM | 477 | CHINA | 10 | PCC-PRD | 2021-12-23 | 2022-12-23 | 5RPM | 101 | 100 | 1 | Okay
[11 empty] | 102 | 100 | 2 | Okay
[11 empty] | 101 | 100 | 1 | Okay
316 | PCC-PRD | MAIN PANNEL | PCC-PR-EQT-21-AMP M 1 | 157 | CHINA | 1 | PCC-PRD | 2021-12-23 | 2022-12-23 | 1A | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
[11 empty] | 0.8 | 0.5 | 0.3 | Okay
317 | PCC-LAB | WEIGHING SALE | PCC-LAB-EQT-08-WS | 109 | CHINA | 10 | PCC-LAB | 2021-12-23 | 2022-12-23 | 10G | 10.3 | 10 | 0.3 | Okay
[11 empty] | 10.6 | 10 | 0.6 | Okay
[11 empty] | 9.8 | 10 | -0.2 | Okay
318 | PCC-LAB | WEIGHING SCALE | PCC-LAB-EQT-007-WS | 110 | CHINA | 10 | PCC-LAB | 2021-12-23 | 2022-12-23 | 10G | 9.9 | 10 | -0.1 | Okay
[11 empty] | 10.6 | 10 | 0.6 | Okay
[11 empty] | 9.8 | 10 | -0.2 | Okay
319 | PCC-LAB | WEIGHING SCALE | PCC-LAB-EQT-11-WS | 111 | CHINA | 10 | PCC-LAB | 2021-12-23 | 2022-12-23 | 10G | 9.9 | 10 | -0.1 | Okay
[11 empty] | 10.6 | 10 | 0.6 | Okay
[11 empty] | 9.8 | 10 | -0.2 | Okay
320 | PCC-LAB | PNEUMATIC MIXER | PCC-LAB-EQT-03-RPM1 | 163 | CHINA | 10 | PCC-LAB | 2021-12-23 | 2022-12-23 | 5RPM | 99 | 100 | -1 | Okay
[11 empty] | 102 | 100 | 2 | Okay
[11 empty] | 101 | 100 | 1 | Okay
321 | PCC-LAB | PNEUMATIC MIXER | PCC-LAB-EQT-03-AMP1 | 162 | CHINA | 1 | PCC-LAB | 2021-12-23 | 2022-12-23 | 1A | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
[11 empty] | 0.8 | 0.5 | 0.3 | Okay
322 | PCC-LAB | BEAD MILL | PCC-LAB-EQT-01-AMP1 | 160 | CHINA | 1 | PCC-LAB | 2021-12-23 | 2022-12-23 | 1A | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
[11 empty] | 0.8 | 0.5 | 0.3 | Okay
323 | PCC-LAB | BEAD MILL | PCC-LAB-EQT-01-RPM1 | 164 | CHINA | 10 | PCC-LAB | 2021-12-23 | 2022-12-23 | 5RPM | 99 | 100 | -1 | Okay
[11 empty] | 102 | 100 | 2 | Okay
[11 empty] | 101 | 100 | 1 | Okay
324 | PCC-LAB | BEAD MILL | PCC-LAB-EQT-02-TG1 | 200 | CHINA | 1 | PCC-LAB | 2021-12-23 | 2022-12-23 | 5 C | 20 | 20 | 0 | Okay
[11 empty] | 40 | 40.3 | -0.3 | Okay
[11 empty] | 55 | 55.1 | -0.1 | Okay
325 | PCC-LAB | BEAD MILL | PCC-LAB-EQT-02-PG1 | 199 | CHINA | 0.1 | PCC-LAB | 2021-12-23 | 2022-12-23 | 0.05 MPA | 0.4 | 0.65 | -0.25 | Okay
[11 empty] | 0.5 | 0.65 | -0.15 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
326 | PCC-LAB | BEAD MILL | PCC-LAB-EQT-02-PG2 | 198 | CHINA | 0.01 | PCC-LAB | 2021-12-23 | 2022-12-23 | 0.05 MPA | 0.4 | 0.65 | -0.25 | Okay
[11 empty] | 0.5 | 0.65 | -0.15 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
327 | PCC-LAB | BEAD MILL | PCC-LAB-EQT-02-PG3 | 201 | CHINA | 0.01 | PCC-LAB | 2021-12-23 | 2022-12-23 | 0.05 MPA | 0.4 | 0.65 | -0.25 | Okay
[11 empty] | 0.5 | 0.65 | -0.15 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
328 | PCC-LAB | BEAD MILL | PCC-LAB-EQT-02-TIMER1 | 372 | CHINA | 1 | PCC-LAB | 2021-12-23 | 2022-12-23 | 0.5Sec | 0.5 | 0.5 | 0 | Okay
[11 empty] | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.8 | 0.5 | 0.3 | Okay
329 | PCC-LAB | BEAD MILL | PCC-LAB-EQT-02-AMP1 | 159 | CHINA | 1 | PCC-LAB | 2021-12-23 | 2022-12-23 | 1A | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
[11 empty] | 0.8 | 0.5 | 0.3 | Okay
330 | PCC-LAB | RK-Printing Proofer | PCC-LAB-EQT-16-RPM1 | 501 | CHINA | 10 | PCC-LAB | 2021-12-23 | 2022-12-23 | 5RPM | 99 | 100 | -1 | Okay
[11 empty] | 102 | 100 | 2 | Okay
[11 empty] | 101 | 100 | 1 | Okay
330 | PCC-LAB | Rub Fastness Tester | PCC-LAB-EQT-17-RPM1 | 502 | CHINA | 10 | PCC-LAB | 2021-12-23 | 2022-12-23 | 5RPM | 99 | 100 | -1 | Okay
[11 empty] | 102 | 100 | 2 | Okay
[11 empty] | 101 | 100 | 1 | Okay
330 | PCC-LAB | MIXER | PCC-LAB-EQT-04-RPM1 | 166 | CHINA | 10 | PCC-LAB | 2021-12-23 | 2022-12-23 | 5RPM | 99 | 100 | -1 | Okay
[11 empty] | 102 | 100 | 2 | Okay
[11 empty] | 101 | 100 | 1 | Okay
331 | PCC-LAB | MIXER | PCC-LAB-EQT-04-AMP1 | 371 | CHINA | 1 | PCC-LAB | 2021-12-23 | 2022-12-23 | 1A | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
[11 empty] | 0.8 | 0.5 | 0.3 | Okay
332 | PCC-LAB | MIXER | PCC-LAB-EQT-05-RPM1 | 167 | CHINA | 10 | PCC-LAB | 2021-12-23 | 2022-12-23 | 5RPM | 103 | 100 | 3 | Okay
[11 empty] | 101 | 100 | 1 | Okay
[11 empty] | 101 | 100 | 1 | Okay
333 | PCC-LAB | MIXER | PCC-LAB-EQT-05-AMP1 | 202 | CHINA | 1 | PCC-LAB | 2021-12-23 | 2022-12-23 | 1A | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
[11 empty] | 0.8 | 0.5 | 0.3 | Okay
334 | PCC-LAB | ELECTRICAL SUPPLY BOX | PCC-LAB-EQT-18-VM1 | 197 | CHINA | 1 | PCC-LAB | 2021-12-23 | 2022-12-23 | 0.5V | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
[11 empty] | 0.8 | 0.5 | 0.3 | Okay
335 | PCC-LAB | HEAT SEALING TESTER | PCC-LAB-EQT-19-TC1 | 168 | CHINA | 1 | PCC-LAB | 2021-12-23 | 2022-12-23 | 5 C | 149 | 150 | -1 | Okay
[11 empty] | 77 | 75 | 2 | Okay
[11 empty] | 1 | 1 | 0 | Okay
336 | PCC-LAB | HEAT SEALING TESTER | PCC-LAB-EQT-19-TC2 | 158 | CHINA | 1 | PCC-LAB | 2021-12-23 | 2022-12-23 | 5 C | 150 | 150 | 0 | Okay
[11 empty] | 74 | 75 | -1 | Okay
[11 empty] | 1 | 1 | 0 | Okay
337 | PCC-LAB | HEAT SEALING TESTER | PCC-LAB-EQT-19-TIMER1 | 377 | CHINA | 1 | PCC-LAB | 2021-12-23 | 2022-12-23 | 0.5Sec | 3 | 3 | 0 | Okay | Okay
[11 empty] | 1.4 | 1.5 | -0.1 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay | Okay
[11 empty] | 0.8 | 0.5 | 0.3 | Okay | Okay
338 | PCC-LAB | HEAT SEALING TESTER | PCC-LAB-EQT-19-PG1 | 376 | CHINA | 0.01 | PCC-LAB | 2021-12-23 | 2022-12-23 | 0.05 MPA | 0.1 | 0.1 | 0 | Okay
[11 empty] | 0.301 | 0.3 | 0.001 | Okay
[11 empty] | 0.59 | 0.6 | -0.01 | Okay
339 | PCC-LAB | AIR COMPRESSOR | PCC-LAB-EQT-19-PS | 456 | CHINA | ON/Off | PCC-PRD | 2021-12-23 | 2022-12-23 | 0.05 MPA | 0.8 | 0.8 | 0 | Okay
[11 empty] | 0.3 | 0.3 | 0 | Okay
[11 empty] | 0.1 | 0.1 | 0 | Okay
340 | PCC-LAB | AIR COMPRESSOR | PCC-LAB-EQT-19-PG1 | 487 | CHINA | 0.01 | PCC-LAB | 2021-12-23 | 2022-12-23 | 0.05 MPA | 0.8 | 0.8 | 0 | Okay
[11 empty] | 0.3 | 0.3 | 0 | Okay
[11 empty] | 0.1 | 0.1 | 0 | Okay
341 | PCC-LAB | HIGH SPEED DISPERSER 0.37 KW | PCC-LAB-EQT--25-AMP | 462 | CHINA | 0.5A | PCC-LAB | 2021-12-23 | 2022-12-23 | 1A | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.8 | 0.5 | 0.3 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
342 | PCC-LAB | HIGH SPEED DISPERSER 0.37 KW | PCC-LAB-EQT-25-RPM | 461 | CHINA | 10 | PCC-LAB | 2021-12-23 | 2022-12-23 | 5RPM | 101 | 100 | 1 | Okay
[11 empty] | 100 | 100 | 0 | Okay
[11 empty] | 101 | 100 | 1 | Okay
343 | PCC-LAB | HIGH SPEED DISPERSER 0.37 KW | PCC-LAB-EQT-26-AMP | 464 | CHINA | 0.5A | PCC-LAB | 2021-12-23 | 2022-12-23 | 1A | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.8 | 0.5 | 0.3 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
344 | PCC-LAB | HIGH SPEED DISPERSER 0.37 KW | PCC-LAB-EQT-26-RPM | 463 | CHINA | 10 | PCC-LAB | 2021-12-23 | 2022-12-23 | 5RPM | 101 | 100 | 1 | Okay
[11 empty] | 100 | 100 | 0 | Okay
[11 empty] | 101 | 100 | 1 | Okay
345 | PCC-LAB | CHILLER | PCC-LAB-EQT-17-TC | 455 | CHINA | 1 | PCC-LAB | 2021-12-23 | 2022-12-23 | 5 C | 50 | 54 | -4 | Okay
[11 empty] | 60 | 62 | -2 | Okay
[11 empty] | 70 | 66 | 4 | Okay
346 | PCC-LAB | BEAD MILL | PCC-LAB-EQT-27-RPM | 481 | CHINA | 10 | PCC-LAB | 2021-12-23 | 2022-12-23 | 5RPM | 101 | 100 | 1 | Okay
[11 empty] | 100 | 100 | 0 | Okay
[11 empty] | 101 | 100 | 1 | Okay
347 | PCC-LAB | BEAD MILL | PCC-LAB-EQT-27-AMP | 482 | CHINA | 0.5A | PCC-LAB | 2021-12-23 | 2022-12-23 | 1A | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.8 | 0.5 | 0.3 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
348 | PCC-LAB | WATER BATH | PCC-LAB-EQT-28-TC | 483 | CHINA | 1 | PCC-LAB | 2021-12-23 | 2022-12-23 | 5 C | 50 | 54 | -4 | Okay
[11 empty] | 60 | 62 | -2 | Okay
[11 empty] | 70 | 66 | 4 | Okay
349 | PCC-LAB | DRYER | PCC-LAB-EQT-29-PG | 488 | CHINA | 0.01 | PCC-LAB | 2021-12-23 | 2022-12-23 | 0.05 MPA | 0.4 | 0.65 | -0.25 | Okay
[11 empty] | 0.5 | 0.65 | -0.15 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
350 | PCC-LAB | DRYER | PCC-LAB-EQT-29-TG | 489 | CHINA | 1 | PCC-LAB | 2021-12-23 | 2022-12-23 | 5 C | 20 | 20 | 0 | Okay
[11 empty] | 40 | 40.3 | -0.3 | Okay
[11 empty] | 55 | 55.1 | -0.1 | Okay
351 | PCC-LAB | COMPRESSOR | PCC-LAB-EQT-33-PG | 490 | CHINA | 0.01 | PCC-LAB | 2021-12-23 | 2022-12-23 | 0.05 MPA | 0.4 | 0.65 | -0.25 | Okay
[11 empty] | 0.5 | 0.65 | -0.15 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
352 | PCC-LAB | COMPRESSOR | PCC-LAB-EQT-33-PS | 491 | CHINA | ON/Off | PCC-PRD | 2021-12-23 | 2022-12-23 | 0.05 MPA | 0.5 | 0.65 | -0.15 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.6 | 0.65 | -0.05 | Okay
353 | PCC-LAB | COMPRESSOR | PCC-LAB-EQT-34-PG | 492 | CHINA | 0.01 | PCC-LAB | 2021-12-23 | 2022-12-23 | 0.05 MPA | 0.4 | 0.65 | -0.25 | Okay
[11 empty] | 0.5 | 0.65 | -0.15 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
354 | PCC-LAB | COMPRESSOR | PCC-LAB-EQT-34-PS | 493 | CHINA | ON/Off | PCC-PRD | 2021-12-23 | 2022-12-23 | 0.05 MPA | 0.5 | 0.65 | -0.15 | Okay
[11 empty] | 0.3 | 0.65 | -0.35 | Okay
[11 empty] | 0.6 | 0.65 | -0.05 | Okay
322 | PCC-LAB | BEAD MILL | PCC-LAB-EQT-40-AMP1 | 601 | CHINA | 1 | PCC-LAB | 2021-12-23 | 2022-12-23 | 1A | 0.3 | 0.5 | -0.2 | Okay
[11 empty] | 0.5 | 0.5 | 0 | Okay
[11 empty] | 0.8 | 0.5 | 0.3 | Okay
323 | PCC-LAB | BEAD MILL | PCC-LAB-EQT-40-RPM1 | 602 | CHINA | 10 | PCC-LAB | 2021-12-23 | 2022-12-23 | 5RPM | 99 | 100 | -1 | Okay
[11 empty] | 102 | 100 | 2 | Okay
[11 empty] | 101 | 100 | 1 | Okay
Note:
Calibrated By: | [11 empty] | Verified By:
Abrivations
1 | TG |  | TEMPRATURE GAUGE
2 | PG |  | PRESSURE GAUGE
3 | AMP M |  | AMPERE METER
4 | TIMER |  | TIMER
5 | TM |  | TEMPRATURE METER
6 | TS |  | TEMPRATUER SENSOR
7 | WS |  | WEIGHING SCALE
8 | RPM M |  | REVOLOUTION PER MINUTE METER
9 | DIG AMP M |  | DIGITAL AMPERE METER
10 | VM |  | VOLT METER
11 | A |  | AMPERE METER
12 | OVEN |  | OVEN
13 | DIG V M |  | DIGITAL VOLT METER
14 | TEMP BEN M |  | TEMPRATURE BENCH METER
15 | LIG B |  | LIGHT BOX
16 | HSV |  | HYDRULIC SOLINOID VALVE
17 | RFM

## Sheet1
Calibrated By: |  |  | Verified By:
Abrivations
1 | TG | TEMPRATURE GAUGE
2 | PG | PRESSURE GAUGE
3 | AMP M | AMPERE METER
4 | TIMER | TIMER
5 | TM | TEMPRATURE METER
6 | TS | TEMPRATUER SENSOR
7 | WS | WEIGHING SCALE
8 | RPM M | REVOLOUTION PER MINUTE METER
9 | DIG AMP M | DIGITAL AMPERE METER
10 | VM | VOLT METER
11 | A | AMPERE METER
12 | OVEN | OVEN
13 | DIG V M | DIGITAL VOLT METER
14 | TEMP BEN M | TEMPRATURE BENCH METER
15 | LIG B | LIGHT BOX
16 | HSV | HYDRULIC SOLINOID VALVE
17 | RFM